## Total
UOT20200101- Spinning Disk Microscope
Scoring Summary
Active Submissions
| Total | A - Submission Requirements | A-1 - Proposal Submission Form | A-2 - Conflict of Interest Declaration | A-3 - Reference Form | B - Technical Submission | B-1 - Experience and Qualifications | B-2 - Spinning Disk Microscope | B-3 - High Speed Imaging Capacity | B-4 - Simultaneous Dual Fluorophore Imaging Solution | B-5 - Photobleaching/Photoconversion Solution | B-6 - Delivery, Installation, Testing, and Training Requirements | B-7 - Warranty and Long-Term Service Agreement | B-8 - Additional Optional Components | C - Financial Submission | C-1 - Pricing | D - Demonstration | D-1 - Demonstration
Supplier | / 100 pts | / 0 pts | Pass/Fail | Pass/Fail | Pass/Fail | / 80 pts | / 5 pts | / 25 pts | / 10 pts | / 15 pts | / 10 pts | / 5 pts | / 5 pts | / 5 pts | / 20 pts | / 20 pts | / 0 pts | Pass/Fail
Nikon Canada Inc | 74.7 | 0 | Pass | Pass | Pass | 56 | 3.667 | 18.33 | 7.333 | 11 | 7.333 | 3 | 3.333 | 2 | 18.7 | 18.7 ($649,819.20) | - | -
Quorum Technologies Inc | 69.33 | 0 | Pass | Pass | Pass | 49.33 | 3 | 15 | 5.333 | 10 | 6.667 | 3 | 2.667 | 3.667 | 20 | 20 ($607,607.00) | - | -
Olympus Canada Inc. | 62.65 | 0 | Pass | Pass | Pass | 43 | 3.333 | 13.33 | 6 | 7 | 5.333 | 2.667 | 2 | 3.333 | 19.65 | 19.65 ($618,477.29) | - | -

## Rank #1
UOT20200101 - Spinning Disk Microscope
Nikon Canada Inc
Scoring Summary
Evaluation Group 1 - Submission Requirements
| A - Submission Requirements | A-1 - Proposal Submission Form | A-2 - Conflict of Interest Declaration | A-3 - Reference Form
Reviewer | / 0 pts | Pass/Fail | Pass/Fail | Pass/Fail
Nathan King | 0 | Pass | Pass | Pass
| Average: | Pass | Pass | Pass
| Consensus: | - | - | -
|  | ↓ | ↓ | ↓
Calculated: | 0 | Pass | Pass | Pass
Evaluation Group 2 - Technical Submission
| Total | B - Technical Submission | B-1 - Experience and Qualifications | B-2 - Spinning Disk Microscope | B-3 - High Speed Imaging Capacity | B-4 - Simultaneous Dual Fluorophore Imaging Solution | B-5 - Photobleaching/Photoconversion Solution | B-6 - Delivery, Installation, Testing, and Training Requirements | B-7 - Warranty and Long-Term Service Agreement | B-8 - Additional Optional Components
Reviewer | / 80 pts | / 80 pts | / 5 pts | / 25 pts | / 10 pts | / 15 pts | / 10 pts | / 5 pts | / 5 pts | / 5 pts
Dr Simple Simon | 54 | 54 | 4 / 5 | 4 / 5 | 4 / 5 | 3 / 5 | 3 / 5 | 3 / 5 | 3 / 5 | 1 / 5
Peter Pan | 55 | 55 | 4 / 5 | 3 / 5 | 3 / 5 | 4 / 5 | 4 / 5 | 3 / 5 | 3 / 5 | 4 / 5
Donald Duck | 62 | 62 | 3 / 5 | 4 / 5 | 4 / 5 | 4 / 5 | 4 / 5 | 3 / 5 | 4 / 5 | 4 / 5
|  | Average: | 3.667 / 5 | 3.667 / 5 | 3.667 / 5 | 3.667 / 5 | 3.667 / 5 | 3 / 5 | 3.333 / 5 | 3 / 5
|  | Consensus: | - | - | - | - | - | - | - | 2 / 5
|  |  | ↓ | ↓ | ↓ | ↓ | ↓ | ↓ | ↓ | ↓
Calculated: | 56 | 56 | 3.667 | 18.33 | 7.333 | 11 | 7.333 | 3 | 3.333 | 2
Evaluation Group 3 - Financial Submission
| Total | C - Financial Submission | C-1 - Pricing
Reviewer | / 20 pts | / 20 pts | / 20 pts
Nathan King | 18.7 | 18.7 | $649,819.20
|  | Average: | $649,819.20
|  | Consensus: | -
↓
Calculated: | 18.7 | 18.7 | 18.7 ($649,819.20)
Evaluation Group 4 - Demonstration
| D - Demonstration | D-1 - Demonstration
Reviewer | / 0 pts | Pass/Fail
Dr Simple Simon | - | -
Peter Pan | - | -
Donald Duck | - | -
| Average: | -
| Consensus: | -
↓
Calculated: | - | -
Scoring Comments
A-1 - Proposal Submission Form - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and signed
A-2 - Conflict of Interest Declaration - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and Signed
A-3 - Reference Form - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete
B-1 - Experience and Qualifications - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 4 / 5 | Well-supported claim(s) | Excellent experience.
Peter Pan | 4 / 5 | Meets or exceeds my expectations | a) Meets expectations. b) Meets/above expectations. Local support team with 2 direct University of Toronto contacts, combined 15 years experience in microscopy systems. c) Meets/above expectations. Proven expertise with sales team in regard to live-cell imaging at the University of Toronto. Experience with the Ti-E microscope system for over a decade. d) Meets/above expectations. Listed contacts with same spinning disk (W1) and inverted microscope. All references are from universities in the USA. The research field/projects are not specified (e.g. live-cell imaging, plant biology, etc.) e) Meets/above expectations. Well organized and manageable work plan for delivery and set-up of microscope. Both immediate and long-term support options are listed. Hierarchical organization of support team well established.
Donald Duck | 3 / 5 | Partially fits desired attributes | - Proponent description requirement met - Does not specifically name any of the employees who will be doing the delverables - Good detail about the employees and description but could mention which person will be doing the deliverables - Adequate description of past experience in delivering the deliverables; adequate descirption describing the knowledg and skills - Good description of how the deliverables will be made - More details in the reference form required to address how the microscopes were used in the past but detailed description of microscope specifications
B-2 - Spinning Disk Microscope - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr Simple Simon | 4 / 5 | Strongly fits desired attribute(s) | Meets criteria for all components. Exceeds criteria for objectives (additional 20x objective included). Exceeds criteria for field of view of microscope and spinning disk. Exceeds criteria for computer (CPU & Storage). Exceeds criteria for software. Details missing for even illumination solution.
Peter Pan | 3 / 5 | Partially fits desired attributes | a-i) Meets expectations, as all filters, xyz movement, and objectives, etc. are motorized through the microscope stand. a-ii) Meets/exceeds expectations. Suitable for standard microscope slide. Additional stage inserts can hold chambered slides, 35-60mm dishes, etc. for additional flexibility during imaging. a-iii) Meets/exceeds expectations, with 2 optical ports with a very high (25mm FOV). a-iv) Meets expectations. a-v) Meets expectations. b) Meets/exceeds expectations, with 4 offered objectives. NA of 60X and 100X objectives higher than requested. c-i) Meets/slightly below expectations. Laser lines will excite outlined fluorophores. Output power sufficient for live cell imaging. Unclear how customizable the laser set up will be in the future, given that the laser launch appears permanently set. c-ii) Meets expectations. d) Meets expectations. Filters and mirrors are sufficient for imaging selected fluorophores and separating CFP/YFP and GFP/RFP pairs. Mirror holders and filter wheels are customizable, although the dichroic mirror holder is limited to 3 mirrors (unclear how the switching of these mirrors is undertaken, as 4 dichroic mirrors are proposed here). Narrow range emission filters will give higher contrast but lower brightness. e-i) Meets/exceeds expectations as all objectives will be sufficient to image multiple plant cells at once. e-ii) Meets/below expectations. The large FOV of the microscope/cameras limits the speed of imaging down to 35fps in 16-bit mode. Despite the slower imaging, this should be sufficient for imaging particles moving >10um/sec. e-iii) Meets expectations. e-iv) Meets expectations, although it is unclear whether there are additional microscope/software features that can correct for this. e-v) Meets expectations. e-vi) Meets expectations. Pinhole size optimal for high magnification, pinhole spacing optimal for reducing cross-talk, and 4000rpm for live-cell imaging. f) Meets/exceeds expectations. Large xy span of stage with very fast drive speed of 25 mm/sec. Z-drive can travel far distances quickly with high resolution/sensitivity. g) Meets expectations. Laser/microscope based system to maintain focus. 700msec correction speed for 5um drift, which is sufficiently fast for using PFS during live-cell imaging. h-i) Meets expectations, and is comparable to EMCCDs. Dynamic range is very high but has a larger read noise, so dim and bright signals can be differentiated more easily (dynamic range) but have potentially higher background noise. h-ii) Meets expectations for 2Mp resolution. h-iii) Meets/exceeds expectations, with an extremely high FOV for all magnifications, suitable for imaging numerous cells at once. h-iv) Meets expectations. High FOV compromises speed of acquisition, but should be sufficient for live-cell imaging. h-v) Meets expectations. Extra large FOV but limited by higher noise levels, higher pixel size (11um), and 2Mp resolution. Features are provided to deconvolute/correct for lower resolution in software. Liquid-based cooling of camera is not included but is upgradable. i) Meets/exceeds expectations. Both spinning disk dependent and independent options for imaging. j) Meets expectations. k) Meets/exceeds expectations. Higher CPU, higher amount of data storage TB, and larger monitor than requested. l-i) Meets expectations. l-ii) Meets expectations. Ability to use previously used/set microscope settings beneficial for new users or users doing repeated experiments. l-iii) Meets/exceeds expectations, with the software being sufficient to perform complex imaging outlined in the RFP. It is unclear whether photoconversion/bleaching capability is also included in these modules. l-iv) Meets expectations. m) Meets expectations. Table and compressor both supplied (in-house air not necessary).
Donald Duck | 4 / 5 | Meets or exceeds my expectations | - really detailed description re: the motorized stage - really detailed description re: stage; includes what other slide sizes and types that can be used - Meets optial ports requirements - Meers DIC, phase contrast, widefiedl flo requirements - Mentions the year of the model as well the accolades it has received - incudes required objective lens including an extra 20x lens; objective lens have higher NA than requested - Laser requirements are met - lasers correpsond to the desired fluorophores - great detail about the emisson fileters and dichroic mirros; lots of room for future filter set installments - provided in great detail about the field of view; lists field of view for each objective lens - Meets and exceeds requirement for high speed imaging - Meets requirements for pinhole size - Meets requirements for even illumination - Meets requirements for two camera ports - Provided all the details for the spinning disk requirements - Provided all the details for the motorized stage - Provided all the details for z-drfit correction - Provided all the detail for the camera; offers deconvolution software; mentioned field of view specifications for each objective lens - Provided good details for BF imaging: even illumination and removal of spinning disk for better BF images - Meets requirements for computer work stations - offers faster processor; more data storage; bigger monitor - Meets requirements for software requirements; provided really good details for novice and advanced users; advanced user software offers 6D imaging - Floating table requirements are met
B-3 - High Speed Imaging Capacity - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 4 / 5 | Meets or exceeds my expectations | Meets criteria.
Peter Pan | 3 / 5 | Partially fits desired attributes | a) Meets/exceeds expectations, with proposed high speed motorized stage and high resolution z-drive. b) Meets expectations, with proposed high speed motorized stage and high resolution z-drive. c) Meets expectations. d) Meets expectations. PFS is microscope-based (not software based). Quick sampling rate and 5um focus drift sufficient for live-cell imaging. e) Meets expectations. Due to large FOV, maximum camera frame rate is slower, but should be sufficient for most live-cell imaging. f) Meets/exceeds expectations, with high GPU and data drives for fast image analysis. g) Meets/exceeds expectations. Dual triggering systems will provide very rapid and seamless live-cell imaging experiments with little delay/feedback. h) Meets expectations. i) Meets expectations. Slower fps is rate-limiting step, so for live-cell imaging a smaller ROI is suggested (although this maybe defeats the purpose of having such a large FOV in the Ti2E system). Also unclear whether another rate limiting step is image acquisition/saving on the computer, which is generally a rate-limiting step for most imaging.
Donald Duck | 4 / 5 | Meets or exceeds my expectations | - Meets requirements for microscope stage movement : speed and range - Meets requirements for SD speed - Meets requirements for speed of light path movement - Meets requirements for z-drift; corrects for vibrations, temp, etc. - Provides details for camera frame rate: frame rate depends on bit-mode - Meets requirements for computer work station; - Explains in great detail about solutions to bypass software latency - High speed imaging can be integrated with software platform - Meets requirements for high-speed imaging capacity; speed can be increased by increasing binning or reducing field of view
B-4 - Simultaneous Dual Fluorophore Imaging Solution - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Medium level of detail in response | Meets criteria, two high quality cameras included.
Peter Pan | 4 / 5 | High level of detail in response | a) Meets expectations. Dual camera and dichroics/filters to separate RFP/GFP and YFP/CFP. b) Meets/exceeds expectations. Large FOV for imaging multiple plant cells simultaneously. c) Meets expectations. Fps sufficient for live cell imaging in 16-bit resolution. d) Meets expectations, with two cameras provided. e) Meets expectations. f) Meets expectations. g) Meets expectations. High quantum efficiency, large FOV, and large dynamic range. However, slightly higher pixel size, read noise, and frame rates. (Higher field of view and sensitivity but reduced speed/resolution and higher background).
Donald Duck | 4 / 5 | Strongly fits desired attribute(s) | - Meets requirements for simultaneous dual fluorophore imaging via the two cameras and types of filters being offered - Meets requirements for the field of view; mentions field of views for each objective lens - Meets requirements for high-speed imaging - Has two cameras and two ports for dual flurorophore imaging and no need for split chip - Software allows for simultaneous dual fluorophore imaging - Provides excellent details for solutions about the simultaneous dual fluorophore imaging
B-5 - Photobleaching/Photoconversion Solution - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr Simple Simon | 3 / 5 | Medium level of detail in response | Meets criteria, but with UV laser all fluorophores will be bleached, not specific ones.
Peter Pan | 4 / 5 | Strongly fits desired attribute(s) | a) Meets expectations. 405 laser to perform both photobleaching/photoactivation. 50mW laser power should be sufficient for photobleaching. b) Meets/exceeds expectations. Ability to personalized region of interest for photomanipulation at very high resolution. c) Meets/exceeds expectations. Very beneficial to have photobleaching and imaging occur in real time/sequentially (rather than have a lag between photobleaching and imaging). Also all laser intensity controlled through software, making it easy to manipulate. d) Meets expectations. e) Meets expectations. f) Meets expectations.
Donald Duck | 4 / 5 | Strongly fits desired attribute(s) | - Meets requirements for photobleaching desired fluorophores - Meets requirements for setting a defined regions; can make defined regions, lines, points, and multiple regions - Mentions all the capacity details for photobleaching/conversion - Can control dwell time and laser output to control photoactivation/conversion; can also do normal imaging at the same time as photobleaching/conversion - Laser is permanently pre-aligned with no maintenance required - Photobleaching/conversion is integrated into the system -
B-6 - Delivery, Installation, Testing, and Training Requirements - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially meets my expectations | Timeline for delivery and installation is 6-14 weeks, slightly longer than criteria, but detailed training plan presented, including followup training.
Peter Pan | 3 / 5 | Medium level of detail in response | a) Meets expectations. b) Meets expectations. More detail about floor loading/weight for crates would have been useful (limited space in building/hallway). c) Meets expectations. d) Meets expectations. e) Meets expectations. f) Meets expectations. g) Meets expectations. h) Meets/below expectations. Would like more specifics about the training available through Nikon - e.g. are there limited hours for training? Does the training have to be performed within a given time frame?, etc.
Donald Duck | 3 / 5 | Mostly complete response | - Delivery time may exceed desired timeframe due to pandemic - Meets requirements for on-site delivery, installation, testing; mentions delivery of packages, space needed, environment needed, etc - Can replace instrument if damaged during delivery - Lists all electrical, data, and environmental conditions required - Meets provincial electrical safety requirements - Meets all installation requirements; can train as many people as needed - Can be delivered at desired location - Provides multiple graduate trainign sessions - Does not mention if it will provide a hardcopy or electronic handbook - Unclear if they will provide training for technical aspects or maintenance
B-7 - Warranty and Long-Term Service Agreement - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially fits desired attributes | Meets criteria.
Peter Pan | 3 / 5 | Partially fits desired attributes | a) Meets/exceeds expectations, with 5 year warranty on mechanical/optical parts. b) Meets expectations. c) Meets expectations. d) Meets expectations. e) Meets expectations. f) Meets expectations. Unclear whether the technician(s) will come from Nikon. g) Meets expectations. h) Meets/below expectations. Software updates are only provided for a year, which could be costly in the long term depending on updates/bug fixes.
Donald Duck | 4 / 5 | Strongly fits desired attribute(s) | - Exceeds 1 year warranty requirement - Meets optional extended 4 year warranty requirement - warranty for defects in product is limited to five years - Technical support is unlimited: meets requirement - Meets expected techincal respsonse time for call and meets remote support services - Does not mention that the technician will be from the same company as the instrument vendor - Meets requirement for technical response time; provides shipping details; provides replacement parts but limited to 8 years - Training provided for technicians - does not supply certificates - Upgrades are free and meet requirements
B-8 - Additional Optional Components - Consensus Score
Name | Consensus Score | Reason | Comments
Nathan King | 2 / 5 | Partially meets my expectations | Provided brief answers without providing the specs or pricing to be considered for lasers and filters. Saying yes, they were able to provide these without the details wasn't complete enough of a response.
B-8 - Additional Optional Components - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr Simple Simon | 1 / 5 | Missing / mismatched attributes | Very few optionals are included, and only the bottom ranked ones. Unclear whether additional laser lines can be easily added.
Peter Pan | 4 / 5 | Strongly fits desired attribute(s) | a) No dual disk offered. This makes low objective imaging less optimal due to pinhole size. b) Meets/below expectations. As mentioned in B2 above, unclear how flexible/easy it is to modify the laser modules and add in additional laser lines. c) Meets expectations. d) Meets/exceeds expectations. Optional 20X objective provided in section B2 above. e) Meets/exceeds expectations. Optional stages already provided in B2 above. f) Meets expectations. g) Meets expectations. h) Meets expectations.
Donald Duck | 4 / 5 | Meets or exceeds my expectations | - Lists all the options or upgrades that the lab may want in the future: emission filters, objectives, laser lines, stage inserts, more widefield epiflo. - Doesn't include a second SD
C-1 - Pricing - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | 18.7 ($649,819.20) | - | Table 1 fees

## Rank #2
UOT20200101- Spinning Disk Microscope
Quorum Technologies Inc
Scoring Summary
Evaluation Group 1 - Submission Requirements
| A - Submission Requirements | A-1 - Proposal Submission Form | A-2 - Conflict of Interest Declaration | A-3 - Reference Form
Reviewer | / 0 pts | Pass/Fail | Pass/Fail | Pass/Fail
Nathan King | 0 | Pass | Pass | Pass
| Average: | Pass | Pass | Pass
| Consensus: | - | - | -
|  | ↓ | ↓ | ↓
Calculated: | 0 | Pass | Pass | Pass
Evaluation Group 2 - Technical Submission
| Total | B - Technical Submission | B-1 - Experience and Qualifications | B-2 - Spinning Disk Microscope | B-3 - High Speed Imaging Capacity | B-4 - Simultaneous Dual Fluorophore Imaging Solution | B-5 - Photobleaching/Photoconversion Solution | B-6 - Delivery, Installation, Testing, and Training Requirements | B-7 - Warranty and Long-Term Service Agreement | B-8 - Additional Optional Components
Reviewer | / 80 pts | / 80 pts | / 5 pts | / 25 pts | / 10 pts | / 15 pts | / 10 pts | / 5 pts | / 5 pts | / 5 pts
Dr Simple Simon | 52 | 52 | 4 / 5 | 3 / 5 | 3 / 5 | 4 / 5 | 3 / 5 | 3 / 5 | 3 / 5 | 3 / 5
Peter Pan | 51 | 51 | 3 / 5 | 3 / 5 | 3 / 5 | 3 / 5 | 4 / 5 | 3 / 5 | 3 / 5 | 4 / 5
Donald Duck | 38 | 38 | 2 / 5 | 2 / 5 | 2 / 5 | 3 / 5 | 3 / 5 | 1 / 5 | 2 / 5 | 4 / 5
|  | Average: | 3 / 5 | 2.667 / 5 | 2.667 / 5 | 3.333 / 5 | 3.333 / 5 | 2.333 / 5 | 2.667 / 5 | 3.667 / 5
|  | Consensus: | - | 3 / 5 | - | - | - | 3 / 5 | - | -
|  |  | ↓ | ↓ | ↓ | ↓ | ↓ | ↓ | ↓ | ↓
Calculated: | 49.33 | 49.33 | 3 | 15 | 5.333 | 10 | 6.667 | 3 | 2.667 | 3.667
Evaluation Group 3 - Financial Submission
| Total | C - Financial Submission | C-1 - Pricing
Reviewer | / 20 pts | / 20 pts | / 20 pts
Nathan King | 20 | 20 | $607,607.00
|  | Average: | $607,607.00
|  | Consensus: | -
↓
Calculated: | 20 | 20 | 20 ($607,607.00)
Evaluation Group 4 - Demonstration
| D - Demonstration | D-1 - Demonstration
Reviewer | / 0 pts | Pass/Fail
Dr Simple Simon | - | -
Peter Pan | - | -
Donald Duck | - | -
| Average: | -
| Consensus: | -
↓
Calculated: | - | -
Scoring Comments
A-1 - Proposal Submission Form - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and Signed
A-2 - Conflict of Interest Declaration - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and Signed
A-3 - Reference Form - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete
B-1 - Experience and Qualifications - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 4 / 5 | Meets or exceeds my expectations | excellent experience.
Peter Pan | 3 / 5 | Mostly complete response | a) Meets expectations. b)At/below expectations. Over 10 years experience for both individuals, adding credibility. However, little detail about the individuals' role/responsibilities in regards to installing/maintenance of the system outlined in this RFP and whether they are locally based in Toronto. The names and responsibilities of subcontractors are also not included. Additional positions are listed in section e), although their roles are not described. c) Meets/above expectations. Outlines the number and location of other spinning disk systems in research labs in Canada. Unclear the specific research field these confocal microscopes are being used for (e.g. plant biology, live-cell imaging, etc.). d) Meets/slightly below expectations. Two references have the same X-light spinning disk as proposed in the RFP. Would like more information in "Nature of Assignment" - e.g. camera type, research area, type of analyses performed on confocal system, etc. e) Below expectations. Did not mention operations manager in section b), although in section e) the operations manager has a central role in the long-term. It is also unclear how frequent the meeting proposed in the "Continuous Service" will be and the nature of these meetings. In particular, the distinction between developing research aims and addressing technical/mechanical problems is not addressed. Finally, the timeline/work plan of delivering, setting up, and training users on the microscope in this RFP is absent.
Donald Duck | 2 / 5 | Medium level of detail in response | - good description of the proponent - roles and responsibilities are not that detailed; missing some information on each employees expertise; missing who will be doing what deliverables - very vague on how their skills and expetise will lead to the delierables - poor reference form: should go into more detail on what the microscopes were used for - detailed description of the deliverables and organizational chart; could mention a tentative timeline for the deliverables. -
B-2 - Spinning Disk Microscope - Consensus Score
Name | Consensus Score | Reason | Comments
Nathan King | 3 / 5 | Partially meets my expectations | Overall meets expectations, provided a strong response for the laser and stage. And some weaker areas included computer/software, and focus.
B-2 - Spinning Disk Microscope - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially fits desired attributes | Meets criteria for all components, but key stats not indicated for focus drift correction system. Exceeds criteria for Lasers and Filters since additional laser lines included, however some fluorophres is sub optimal with the lasers indicated (excitation of mNeon is only 40% with 470 laser eGFP is only 78%). Exceeds criteria for speed and accuracy of z-drive; XY drive meets criteria, though some info missing. Exceeds criteria for camera since there are additional benefits from three camera imaging modes (high speed, medium, high sensitivity). Exceeds criteria for software (has added features for protecting advanced settings from novice users).
Peter Pan | 3 / 5 | Medium level of detail in response | a-i) Meets expectations. a-ii) Meets expectations. More detail would be nice - dimensions of holder, whether adaptors are required/available, etc. a-iii--a-v) Meets expectations. b-i) Meets expectations. b-ii) Meets/exceeds expectations. NA higher than requested. b-iii) Meets expectations. c-i) Above expectations. Numerous additional laser lines to optimize imaging of different fluorophores. Laser power/output is way more than sufficient for imaging. Also options to include 488 and 577 lasers in the future. c-ii) Meets expectations. d) Meets expectations. Dichroics and emission filters are sufficient for separating fluorophores for dual cameras. Does not specifically address whether these filter wheels are further customizable. Also unclear why both the 402/468/555/640 and 402/488/577/640 quad band exciters/emitters are necessary. Emitter bandwidth mostly consistent at small distance across filters - higher sensitivity but lower brightness. e-i) Meets/above expectations for imaging numerous cells at 100X. e-ii) Below expectations. Would have preferred specifics of microscope/spinning disk/camera that would allow rapid live-cell imaging, as the demonstration was very variable in terms of fluorescent signal for the plant lines, type of organelles viewed, lasers used, etc. e-iii) Meets expectations. e-iv) Meets expectations. Although it is unclear whether there are any corrective measures (either on microscope/software) to account for this. e-v) Meets expectations. e-vi) Meets/above expectations. Spinning disk rotation speed extremely fast, and there are options for alternative pinhole size for imaging at lower magnification/in thicker samples. However, FOV for spinning disk (25mm) is much higher than FOV of camera, so imaging FOV will ultimately be limited by the camera, not the spinning disk. f) Meets/below expectations. Volocity of xy movement is not very rapid at 7. Z-drive has large travel range but also an increased minimum resolution at 2.2nm.mm/sec (potentially decreased resolution in Z-stacks). g) Meets/above expectations, as there are two independent autofocus systems. It would have been useful to include an estimate of response times, as this could play a major factor during live-cell imaging. h-i) Meets/exceeds expectations. Very low read noise, thus minimizing background noise from imaging. Also comparable dynamic range to EMCCD. Low pixel size will ensure high resolution images. h-ii) Meets/exceeds expectations. Resolution (Mp) much higher than requested, thus providing higher resolution in imaging. h-iii) Meets expectations. h-iv) Meets/exceeds expectations and will be sufficiently fast for live-cell imaging. h-v) Meets/above expectations. Could not find attached information form. Resolution sentence confusing: 65nm/pixel at 100X and 103nm/pixel at 63X? i) Meets/above expectations, as brightfield images can be taken both with spinning disk or without. j) Meets expectations. k) Meets expectations in that the Proponent states that the computer specifications will be complied with. It is unclear whether the computer workstation specifications will be exactly as listed in the RFP. l-i) Unclear whether simultaneous imaging, FRAP/photoconversion, and other imaging processes are all integrated into MetaMorph. l-ii) Meets expectations. Does not specify whether programs/experiments can be saved for future use, which is an important parameter for new microscopy users. l-iii) Meets/below expectations. Unclear whether the menus need to be set up by the Proponent or whether the interface can be adjusted by the users for a given experiment. l-iv) Meets expectations. Similar to part iii), does not address the degree of autonomy the microscope users have in deciding whether to be in "basic" or "advanced" mode. Also it is unclear how are certain functions are restricted (e.g. by a administrator function). m) Meets/above expectations.
Donald Duck | 2 / 5 | Partially fits desired attributes | - Meets requirements for motorized stage - Vague description about the size of slides that the stage insert can accommodate - Meets requirements for optical ports - Meets requirements for optical upgrades - Meets requirements for microscopy model year - Meets objective lens requirements - Meets requirements for laser lines to image fluorophores; goes into more detail about the output power at the fiber tip and output source - Meets requirements for laser fibre alignments - Detailed description of optics and emission filters but unclear about the number of positions that exists in each filter wheel - Meets requirements for FOV for all objective lens - Meets requirements for high speed imaging - Meets requirements for disk characteristics for high res/magnified imaging - Provides detailed information about the even illumination field - Meets requirements for 2 camera ports - Meets requirements for SD details; provides further info about the advantages of their SD - Meets requirements for description of motorized stage; mentions weight capacity - Partially meets requirements for z-drift correction system: does not mention accuracy or precision - Partially meets requirements for camera; does not mention FOV for other objective lens, - Meets camera requirements - Does not provide any details about the computer work station - Partially meets requirements for software details; vague info on the software for advanced users
B-3 - High Speed Imaging Capacity - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially supported claim(s) | Exceeds criteria due to "Fast mode" of camera and its small pixel size allows binning (though this doesn't increase speed, it increases sensitivity in fast mode), fast mode for changing light path and for triggering, but details lacking on rate limiting step.
Peter Pan | 3 / 5 | Medium level of detail in response | a) Meets/below expectations. XY travel range is large, but slower movement at 7mm/sec (which may impede live-cell imaging). Z-drive has long travel range but slightly higher minimum resolution. b) Meets/exceeds expectations and will be more than sufficient for live-cell imaging. c) Meets/exceeds expectations and will be more than sufficient for live-cell imaging. d) Meets/above expectations, as there are two systems. However, the response times for these feedback modes were not listed, so it is unclear whether this would slow down the rate of live-cell imaging. e) Meets/exceeds expectations and will sufficiently fast for live-cell imaging. f) Meets/below expectations. As the particular specifications for computer workstation in section B.2k were not given, the speed of the computer and saving/viewing the image is unknown. Often saving imaging to the computer is a rate-limiting step for experiments, so it would have been useful to further discuss this. g) Below expectations, as the specific of this trigger system is not elaborated on. Is it part of the microscope or a separate controller? h) Meets expectations. i) Meets expectations. FOV and fps will be more than sufficient for live-cell imaging. Did not include rate-limiting step for imaging/processing. Also, could not find attached information?
Donald Duck | 2 / 5 | Partially meets my expectations | - Meets requirements for stage speed, movement, and accuracy; mentions weight-load of stage - Meets SD speed requirement - Detailed response for changing filter wheels - mentions slow and fast modes - Two feedback modes for speed of real-time focus: only mentions speed for one of these modes and not the other - Meets camera frame rate requirements - Meets computer requirements - Meets requirements for trigger-based control - Meets High speed imaging via software requirements - Mentions imaging rate and solutions to decrease imaging time; does not mention the rate limiting step of imaging
B-4 - Simultaneous Dual Fluorophore Imaging Solution - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 4 / 5 | Meets or exceeds my expectations | Meets criteria, two high quality cameras included.
Peter Pan | 3 / 5 | Mostly complete response | a) Meets expectations. Dichroics and emission filters listed in B2 are sufficient to separate CFP/YFP and GFP/RFP. b) Meets/exceeds expectations, with a large FOV for imaging. c) Meets/below expectations. See section Be-ii) above. The demonstration could be very variable in terms of fluorescent signal for the plant lines, type of organelles viewed, lasers used, etc. However, the maximum proposed for the fps should be sufficient for live-cell imaging at 10um/sec. d) Meets expectations. e) Meets expectations. f) Meets/slightly below expectations. Unclear how the custom journals are integrated into the Metamorph software and their ease of use for regular dual camera imaging. g) Meets/above expectations. Dynamic range and noise characteristics are optimal for high sensitivity imaging, and the small pixel size will provide high resolution for imaging fluorophores of different brightness.
Donald Duck | 3 / 5 | Mostly complete response | - Meets requirements for dual imaging of fluorophore pairs - Detailed response for field of view: Mentions field of view for all objectives - Meets requirements for high-speed imaging - Meets camera requirements - Mentions that software can perform dual imaging of fluorophore pairs - Meets requirements for Simultaneous Dual Fluorophore Imaging Solutions; mentions different camera modes of imaging depending on the experiment; mentions ability to modify camera for different fluorophore pairs
B-5 - Photobleaching/Photoconversion Solution - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially fits desired attributes | Meets most criteria, but with UV laser all fluorophores will be bleached, not specific ones, not fully integrated into imaging software and lasers will require alignment correction every few years.
Peter Pan | 4 / 5 | Meets or exceeds my expectations | a) Meets expectations. 405 laser at 100mW should be sufficient to bleach all fluorophores for FRAP experiments. b) Meets/above expectations. Many features to fine-tune region(s) of interest for photobleaching. Very high scan speed. Would have been good to include spot-size for 100X objective, where most of the FRAP experiments will likely be performed on. Does not mention delay time (if any) between photomanipulation and imaging, which is a major factor for live-cell imaging. c) Meets expectations. d) Meets expectations. Meets/above expectations. Many features to fine-tune region(s) of interest for photoactivation. e) Meets expectations. f) Meets expectations. Two softwares are required for photobleaching/photoconversion, which may make it more difficult for novice users to use.
Donald Duck | 3 / 5 | Mostly complete response | - Meets requirements for photobleaching/conversion desired fluorophores - Meets requirements for ROI selection; mentions different types of ROIs and includes diffraction limits - Meets requirements for laser power and dwell time - Meets requirements for low laser dose experiments; goes into detail about the features associated with this procedure - Meets requirements for permanently pre-aligned lasers; mentions. that it wouldn't require maintenance for several years - Meets software requirements
B-6 - Delivery, Installation, Testing, and Training Requirements - Consensus Score
Name | Consensus Score | Reason | Comments
Nathan King | 3 / 5 | Partially meets my expectations | Overall met the expectations - See Heather and Natalie's response.
B-6 - Delivery, Installation, Testing, and Training Requirements - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Smple Simon | 3 / 5 | Partially meets my expectations | Timeline for delivery and installation is 8-12 weeks, slightly longer than criteria.
Peter Pan | 3 / 5 | Medium level of detail in response | a) Meets/slightly below expectations, although the estimated delivery may be later than the 8 weeks outlined in RFP. b) Meets/below expectations. Unclear the specifics of how delivery/set-up will be coordinated with the University and the estimated time frame for this. c) Meets expectations. d) Meets/slightly below expectations. Does not indicate environmental conditions of the room (e.g. temperature). e) Meets expectations. f) Meets expectations. g) Meets expectations. h) All aspects for training and equipment use are acknowledged and complied with, although there is no detail provided about the number of days/hours provided for training, information provided for usage/care of the microscope.
Donald Duck | 1 / 5 | Poor level of detail in response | - Meets requirements for delivery time - Meets delivery requirements - Meets electrical requirements specifications including surge protection, type of outlets needed; does not list any environment conditions (size of room, size of table for computer) - CSA approved - Does not mention any details about the installation, training, maintenance, etc. -
B-7 - Warranty and Long-Term Service Agreement - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Mostly complete response | Meets criteria.
Peter Pan | 3 / 5 | Mostly complete response | a) Meets/above expectations, as 2 year warranty included. b) Meets expectations. c) Meets expectations. d) Meets expectations. Unspecified whether this is under the 2 year warranty only, or for the lifetime of the microscope. e) Meets expectations. Unspecified whether this is under the 2 year warranty only, or for the lifetime of the microscope. f) Meets expectations. Unclear whether 24 hour response time is within business hours or also within weekends/holidays. In addition, it is unclear whether the 10 year replacement part availability is following discontinuation of the part? g) Meets expectations. h) Meets expectations. Unspecified whether this is under the 2 year warranty only, or for the lifetime of the microscope.
Donald Duck | 2 / 5 | Mostly complete response | - Breaks down the specifics of the warranty including laser hours, number of maintenance visits, coverage of maintenance visits - Does not meet the requirement for a 4 year warranty extension - Does not mention unconditional warranty due to damage from delivery - Meets requirements for details on the technical support and call back time - Meets requirements for technical support response times, availability of replacement parts, and details for shipping -Does not mention training of service personnel - Meets requirements for free software and firmware updates and maintenance
B-8 - Additional Optional Components - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially meets my expectations | Many lasers and filters already included in microscope setup, but including near IR laser would replace 528 and 540 lasers with 730.
Peter Pan | 4 / 5 | Strongly fits desired attribute(s) | a) Meets/slightly below expectations. Would have expected a smaller pinhole size for low magnification imaging (e.g. 25um), as 40um spacing would be optimal for higher objectives. b) Meets/exceeds expectations, as most proposed laser lines are already included. c) Meets expectations. d) Meets expectations. e) Unclear whether this is a new stage or an adapter for the universal stage outlined in section B. Also, unsuitable for larger samples such as multiwell plates. f) Meets/exceeds expectations. g) Meets expectations. h) Meets expectations.
Donald Duck | 4 / 5 | Strongly fits desired attribute(s) | - Meets SD requirement - Meets additional lasers, filter sets, objectives, stage inserts, and DIC components - Describes in better detail the options for stage insert - Describes in better detail the DIC options that are compatible with the objectives
C-1 - Pricing - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | 20 ($607,607.00) | - | Table 1 Fees - Summation of total column completed by procurement as no total was provided.

## Rank #3
UOT20200101 - Spinning Disk Microscope
Olympus Canada Inc.
Scoring Summary
Evaluation Group 1 - Submission Requirements
| A - Submission Requirements | A-1 - Proposal Submission Form | A-2 - Conflict of Interest Declaration | A-3 - Reference Form
Reviewer | / 0 pts | Pass/Fail | Pass/Fail | Pass/Fail
Nathan King | 0 | Pass | Pass | Pass
| Average: | Pass | Pass | Pass
| Consensus: | - | - | -
|  | ↓ | ↓ | ↓
Calculated: | 0 | Pass | Pass | Pass
Evaluation Group 2 - Technical Submission
| Total | B - Technical Submission | B-1 - Experience and Qualifications | B-2 - Spinning Disk Microscope | B-3 - High Speed Imaging Capacity | B-4 - Simultaneous Dual Fluorophore Imaging Solution | B-5 - Photobleaching/Photoconversion Solution | B-6 - Delivery, Installation, Testing, and Training Requirements | B-7 - Warranty and Long-Term Service Agreement | B-8 - Additional Optional Components
Reviewer | / 80 pts | / 80 pts | / 5 pts | / 25 pts | / 10 pts | / 15 pts | / 10 pts | / 5 pts | / 5 pts | / 5 pts
Dr. Simple Simon | 34 | 34 | 3 / 5 | 2 / 5 | 2 / 5 | 2 / 5 | 3 / 5 | 2 / 5 | 0 / 5 | 3 / 5
Peter Pan | 53 | 53 | 4 / 5 | 3 / 5 | 4 / 5 | 3 / 5 | 3 / 5 | 3 / 5 | 4 / 5 | 4 / 5
Donald Duck | 43 | 43 | 3 / 5 | 3 / 5 | 3 / 5 | 2 / 5 | 2 / 5 | 3 / 5 | 3 / 5 | 3 / 5
|  | Average: | 3.333 / 5 | 2.667 / 5 | 3 / 5 | 2.333 / 5 | 2.667 / 5 | 2.667 / 5 | 2.333 / 5 | 3.333 / 5
|  | Consensus: | - | - | - | - | - | - | 2 / 5 | -
|  |  | ↓ | ↓ | ↓ | ↓ | ↓ | ↓ | ↓ | ↓
Calculated: | 43 | 43 | 3.333 | 13.33 | 6 | 7 | 5.333 | 2.667 | 2 | 3.333
Evaluation Group 3 - Financial Submission
| Total | C - Financial Submission | C-1 - Pricing
Reviewer | / 20 pts | / 20 pts | / 20 pts
Dr. Simple Simon | 19.65 | 19.65 | $618,477.29
|  | Average: | $618,477.29
|  | Consensus: | -
↓
Calculated: | 19.65 | 19.65 | 19.65 ($618,477.29)
Evaluation Group 4 - Demonstration
| D - Demonstration | D-1 - Demonstration
Reviewer | / 0 pts | Pass/Fail
Dr. Simple Simon | - | -
Peter Pan | - | -
1111 | - | -
| Average: | -
| Consensus: | -
↓
Calculated: | - | -
Scoring Comments
A-1 - Proposal Submission Form - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and Signed
A-2 - Conflict of Interest Declaration - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and Signed
A-3 - Reference Form - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | Pass | Meets the requirement(s) | Complete and signed
B-1 - Experience and Qualifications - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially supported claim(s) | Vague description of previous experience.
Peter Pan | 4 / 5 | Meets or exceeds my expectations | a) Meets expectations. b) Meets expectations. Local Toronto-based team to address urgent/non-urgent questions. Team is specific towards spinning disk microscopes. c) Meets expectations, although a little vague overall in experience with addressing the requirements of this RFP (e.g. unclear how many IXplore Spin spinning disks/cellFRAP systems are currently in use for plant live-cell imaging, in which countries, etc.). d) Meets expectations. It's good to see a Canadian university (in Calgary) in addition to two universities from the USA. Includes a plant-related field. Few/no details included about camera used for each of these systems. e) Slightly below/meets expectations. The specifics of the work plan are a little vague (e.g. how long for set-up of the microscope). The RFP mentions that they intend to maintain a long term relationship with the University, although it is unclear from other sections of B.1 whether the Proponent has previous connections with the University of Toronto specifically. In addition, it is unclear how the 5 support tiers are hierarchically organized. It would have been nice to include a diagram/chart for visualization.
Donald Duck | 3 / 5 | Partially meets my expectations | - Good description of the employees, their roles/responsibilities, - Could have a better explanation on the deliverable details - Needed an organizational chart showing its working relationship with the university - Very detailed reference form
B-2 - Spinning Disk Microscope - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 2 / 5 | Partially meets my expectations | Meets criteria for most components, but does not meet them for lasers & filters (GFP laser not included, laser lifetime not indicated, information about alignment is vague, doubled up filter set for CFP is inappropriate, laser power indicated, but unclear whether this is at source or fibre tip). Does not meet criteria for spinning disk since key details missing for pinhole crosstalk, pinhole spacing, and even illumination solution. Does not meet criteria for xy drive of focus drift correction system since details not provided. Key details not provided regarding software (no significant details provided on what they mean by "user friendly" or "complex…image acquisition", no discussion of whether all camera features (e.g. high sensitivity, high speed) are integrated into software). Exceeds criteria for objectives (excellent NA).
Peter Pan | 3 / 5 | Partially fits desired attributes | a) Meets expectations. For a-v), it would be useful to see the release dates of the original models/software versions (e.g. original release of CellSens Dimension, then release date of v.3.1) to see how these modules are updated over time. b) Meets/above expectations. NA of all 3 objectives higher than requested. c-i) Below expectations. Laser lines have sufficient power for imaging purposes. However, proposed laser lines (e.g. 448) not optimal for excitation of GFP. But sufficient for other proposed fluorophores. Laser lines are customizable (i.e. can get 2 additional laser lines). c-ii) Meets/above expectations, with no need for alignment by the user, and easy replacement. d) Meets/slightly below expectations. Dichroics/emission filters should be suitable for separating CFP/YFP and GFP/RFP, although single 447nm emission filter not optimal for CFP emission(?). Bandwidth range differs greatly between emission filters, which allows higher brightness but potentially more noise/less contrast in images. e-i) Meets/exceeds expectations, giving a very large FOV for 100X. Would have been nice to include details about FOV in other objectives. e-ii) Meets/exceeds expectations, giving a high fps for live-cell imaging at full frame. e-iii) Meets expectations. e-iv) Meets/above expectations, as there is a microscope-based system to correct for uneven illumination. e-v) Meets expectations. e-vi) Meets expectations. f) Meets expectations. Very fast movement for xy stage, but did not mention dimensions of x-y movement. Z-stage has large distance with high precision. g) Meets expectations. h-i) Very low read noise, thus minimizing background noise from imaging. Dynamic range comparable to EMCCDs. h-ii) Resolution (Mp) much higher than requested, thus providing higher resolution in imaging. h-iii) Meets/exceeds expectations. Frame rate more than sufficient for live cell imaging. h-iv) Meets/exceeds expectations. Frame rate more than sufficient for live cell imaging. h-v) Meets expectations. i) Meets/slightly below expectations. BF image goes through spinning disk, without the option of imaging without the spinning disk. j) Meets expectations. k) Meets expectations, and has greater SSD storage, monitor size, and CPU than requested. l) Meets expectations. m) Meets expectations. University may be responsible for supplying a compressor, if in-house air is not sufficient.
Donald Duck | 3 / 5 | Mostly complete response | - Meets requirements for motorized control fo microscope: mentions all motorized components - Meets requirements for stage insert size - Meets requirements for two optica ports; mentions field of view size - Meets requirements for optical upgrades - Meets requirements for model years for all microscope components including software - Meets requirements for objective lenses - Meets requirements for types of lasers lines and power, optical fibre alignment - Meets requirements for optics and emission filters - Meets requirements for field of view but only mentions FOV for 100x objective - Meets requirements for high-speed imaging; mentions ways to increase speed imaging - Meets requirements for pinhole size and disk speed - Meets requirements for even illumination - Meets requirements for two camera ports - Meets requirements for pinhole size, optical characteristics, frame rates, etc. - Meets requirements for motorized stage - Requirements for Hardware for z-drift correction partially met; does not mention precision - Camera requirements partially met; does not mention FOV for other objectives; does not mention any cooling information - Meets requirements for BF imaging - Does not go into detail about the optical, mechanical, data and electrical connections - Meets computer workstation requirements - Meets requirements for software platform requirements
B-3 - High Speed Imaging Capacity - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 2 / 5 | Partially supported claim(s) | Likely meets criteria, but difficult to judge since details not provided on rate-limiting step.
Peter Pan | 4 / 5 | Strongly fits desired attribute(s) | a) Meets/exceeds expectations. Very fast tracking rate of XYZ with high resolution. b) Meets/exceeds expectations; sufficient speeds for live-cell imaging. c) Meets expectations. d) Meets expectations. e) Meets/exceeds expectations, with a very fast fps for live-cell imaging. f) Meets expectations. g) Meets expectations. Triggering device is a separate controller from microscope/software. h) Meets expectations. i) Meets expectations and should be sufficiently fast for live-cell imaging.
Donald Duck | 3 / 5 | Mostly complete response | - Meets desired requirement for imaging high speed particles - Meets requirements for speed of stage in x, y, z positions - Meets requirements for SD speed - Meets requirements for speed of changing filters/light paths - Meets requirements for speed of z-drift correction - Meets requirements for max camera frame rate; mentions frame size - Mentions components of computer workstation to help high-speed imaging - Able to bypass software - High speed capacity is integrated into the software - Mentions max imaging rate, rate-limiting step and solutions to increase imaging time -
B-4 - Simultaneous Dual Fluorophore Imaging Solution - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 2 / 5 | Partially fits desired attributes | Two high quality cameras included, but does not meet criteria for GFP-RFP imaging since laser lines and filter sets are not optimal.
Peter Pan | 3 / 5 | Partially fits desired attributes | a) Meets expectations. Dichroics listed in part B2 sufficient to separate GFP/RFP and CFP/YFP, although the 4 proposed laser lines in section B2 are not optimal for exciting GFP (~488nm). Single emission filter for CFP is also not optimal for emission wavelength. b) Meets/exceeds expectations, with large FOV for imaging multiple cells at 100X. c) Meets expectations. d) Meets expectations. e) Meets expectations. f) Above expectations, as you can theoretically simultaneously image 4 fluorophores (provided additional beam splitters are purchased). g) Meets expectations.
Donald Duck | 2 / 5 | Mostly complete response | - Meets requirements for simultaneous dual fluorophore imaging - Only mentions field of view for 100x objective lens - Does not mention high-speed imaging with respect to 16 bit resolution - Meets camera requirements- 2 provided - simultaneous dual fluorophore imaging integrated in software - Does not mention the specifications and details for simultaneous dual fluorophore imaging
B-5 - Photobleaching/Photoconversion Solution - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Partially fits desired attributes | Meets most criteria (exceeds with regards to FRAP laser lines), but key details not indicated (e.g. field size, scan speed) and lasers will require alignment correction every few years.
Peter Pan | 3 / 5 | Partially fits desired attributes | a) Above expectations. Provides two laser (405 and 561), providing additional specificity for bleaching different fluorophores. Although it is unclear why the 405 laser is not sufficient to bleach all fluorophores (as is common for most photobleaching experiments). Intensity/power of lasers provided in part c) and is sufficient for bleaching. b) Meets/below expectations. The slight delay in photobleaching and image acquisition could be detrimental for live-cell imaging. Galvo has high speed (2kHz). c) Meets expectations. d) Meets/below expectations. Has a solution for photoconversion of fluorophores, although there is little detail on the manual aspect of changing the filter of the high powered lasers; i.e. will novice microscopy users be able to do this? e) Meets expectations. f) Meets expectations. The manual aspect of setting up FRAP is not well explained [also see part d) above]. The fact that some of the FRAP set-up/acquisition is manual may not be optimal for microscope users that have a broad range of skills.
Donald Duck | 2 / 5 | Mostly complete response | - Meets most requirements for photobleaching/conversion for desired fluorophores; doesn't mention if there's a photoconversion capability for RFP/mCherry fluorophores - Meets custom user-defined regions; doesn't indicate the scanfield size - Meets requirements for control over laser power and dwell time; indicates output power - Lasers can do low and high power dosage - Lasers are permanently pre-aligned - Photobleaching/conversion control and calibration is done through software
B-6 - Delivery, Installation, Testing, and Training Requirements - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 2 / 5 | Partially supported claim(s) | timeline for delivery is 6-8 weeks, slightly shorter than criteria, but timeline for installation not indicated.
Peter Pan | 3 / 5 | Partially fits desired attributes | a) Meets/exceeds expectations, and may be faster than the 8 weeks. b) Meets/slightly below expectations. It may not be desirable to have microscope components delivered without Olympus staff present to receive it (i.e. the University has full responsibility for sorting delivery and transfer to the room). c) Meets expectations. d) Meets expectations. 13 outlets is a high number, even with surge-protected power bars. e) Meets expectations. f) Meets expectations. g) Meets expectations. h-i) Meets/below expectations. 24 hours over the course of 2 years does not seem like sufficient time to train multiple members of the Department for all the specialized functions of the microscope/software. h-ii) Meets/slightly below expectations. Does not specify whether software usage (besides basics) will be provided in training outlined in part h-i) or whether information will be provided in manuals instead. h-iii) Meets expectations. h-iv) Meets expectations.
Donald Duck | 3 / 5 | Partially meets my expectations | - Meets requirements for desired delivery time - Meets requirements for delivery, installation, and testing; does not mention the details for - Goes into great detail about the space, electrical, environmental requirements; mentions required temp of the room, space needed for computer, microscope; describes requirements for floating table - Can be delivered to requested address - CSA approved - On-site training is limited to 24 hours spread out through two years; mentions training large and small groups of people; offers more in-depth training re: details of the system. - Meets requirements for providing documentation for using system, technical aspects, maintenance, etc
B-7 - Warranty and Long-Term Service Agreement - Consensus Score
Name | Consensus Score | Reason | Comments
Nathan King | 2 / 5 | Mostly complete response | Standard warranty stated that warranty only covered Olympus parts and others to be managed by other manufacturers. Unclear on what was included in extended warranty. Didn't include electrical items in standard warranty.
B-7 - Warranty and Long-Term Service Agreement - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 0 / 5 | Missing / mismatched attributes | Does not meet criteria since under "Not covered" it says "products not manufactured by Olympus and/or not bearing the "OLYMPUS" brand label " (i.e. most of the system), only "minor upgrades and security patches for the software and firmware needed for maintaining system functionality are included."
Peter Pan | 4 / 5 | Strongly fits desired attribute(s) | a) Meets/above expectations, due to the 5 year warranty on optical and mechanical components. b) Meets expectations. c) Meets expectations. d) Meets expectations. e) Meets expectations. Unclear whether the 24 hour response time is within the business week (versus off-hours/weekend/holidays, etc.). f) Meets/below expectations. It is good to have representatives in the GTA for easy access to technical support. The description of the standard shipping practices and long-term availability of different components (e.g. after discontinuation) are lacking. g) Meets expectations. h) Slightly below expectations. It is unclear what "minor upgrades" refers to and whether more major upgrades would require additional cost. Also unclear whether this falls under the warranty or for the life-time of the microscope.
Donald Duck | 3 / 5 | Mostly complete response | - Surpasses warranty requirements: 5 year warranty for optical and mechanical components and 1 year warranty for whole system - Includes optional warranty between 1-4 years; indicates longer warranties are cheaper - Guarantees that microscope won't be defective during installation; if they are, they can be easily be replaced free of charge - Meets technical support requirements - Meets technical support response times; available through email, phone, and remote services - Meets response time for technical services; has technicians within GTA that can fix components who work for Olympus; warehouse is near by for replacements; good detail on the technical service structure - Technical service personnel are trained; did not indicate specifics on the quality or amount or certifications - Software updates and upgrades are free -
B-8 - Additional Optional Components - Reviewer Scores
Reviewer | Score | Reason | Comments
Dr. Simple Simon | 3 / 5 | Medium level of detail in response | optional items all included, but unclear whether the second disk affects the base price of the microscope since removing the single disk W1 and replacing with the dual disk W1, essential details not included for filter sets or lasers.
Peter Pan | 4 / 5 | Meets or exceeds my expectations | a) Above expectations. Additional spinning disk with different pinhole size for use with lower objectives. b-d) Meets/above expectations in providing optional options for purchase. e) Meets expectations, although it is unclear whether this is the only other option for stage adaptors. f-h) Meets expectations.
Donald Duck | 3 / 5 | Mostly complete response | - Lists SD, laser lines, objective lens, stage inserts, wide-field flo components that are options for upgrades but are not included with the purchase
C-1 - Pricing - Reviewer Scores
Reviewer | Score | Reason | Comments
Nathan King | 19.65 ($618,477.29) | - | Table 1 fees